[image: image1.jpg]



the hashemite kingdom of jordan
the jordanian electric power company limited

Specification JEP 24/2026
tender document for

infrared thermal camera 
	Tenderer's Name:
	_____________________

	Address:
	_____________________

_____________________

	Closing Date:
	_____________________

	Closing Time:
	_____________________


may 2026
THE JORDANIAN ELECTRIC POWER COMPANY LIMITED

AMMAN   – JORDAN
	Telephone 962 6 5503600
P.O. BOX 618

ZIP CODE 11118
	____________
	CABLE: ELECTRICO

FAX 962 6 5503618
specs@jepco.com.jo


SPECIFICATION JEP 24/2026
infrared thermal camera 

General specification
Instruments shall be used mainly for predictive maintenance purposes for JEPCO over head lines electrical network, cable and transformers.
Instruments shall be easy to use with ergonomic and rugged design and with light weight; also it should be design to withstand shocks, vibration, dust & water.
The infrared instruments shall have weather proofed housing to keep the thermometer during off-use.

item 1: Infrared imaging camera 
High-resolution handheld thermal camera for Overhead Line Patrolling, it should be design so that you can look at the target by carrying the camera and pointing on target and then viewing the image on waterproof LCD display.
The buttons and controls should be conveniently located, the sign and symbol on the buttons should be common sense designed, and also all letters shall be in English with suitable size.

	Description 
	Item 1

	Application 
	Overhead line use

For more than 100 meters 

	Thermal resolution 
	640X480

	Display Size
	Min. 3.5" touch screen

	Visible Light Camera
	Min. 4Mp

	Accuracy 
	± 2%

	Temperature Measurement Range
	-20 to 500ºc

	Frame rate
	50HZ or 60HZ

	Focusing 
	Manual and automatic 

	Zoom
	Min. 8X

	Laser pointer 
	Required

	USB interface 
	Required

	Wireless connectivity
	Wi-Fi & Bluetooth

	GPS 
	Is preferable

	Operating conditions 
	-10º C to 60º C with humidity from 10% to 95%.


The field of view shall be of suitable range with the application that the camera to be used.
The LCD display should show the status of battery, storage media and communication status.

The file format for thermal and visual images should be standard JPEG.

The camera shall be supplied with spare rechargeable battery; the operating time for battery shall be more than three hours.

The battery shall be rechargeable and field replaced, in which the camera should be supplied with cable with stud in order to be charged from car at 12V, and also it should be supplied with two (AC adapter 230 V AC 50 Hz).
The camera should shut down automatically with sleep mode for battery saving purpose.

Scope of work:

The works shall comprise the design, manufacture, works testing, supply and delivery to Queen Alia international Airport- Jordan, the Hashemite Kingdom of Jordan of the following:
	Item
	Description
	Qty

	1
	Infrared imaging camera
	8


Important notes:
· The Tenderer shall submit his offer in two forms; one hard copy and one soft copy on "CD" as a PDF. file in a closed envelope. Both copies shall be exactly identical and shall include complete documentation necessary for evaluation such as technical and commercial information, schedules and drawings.

· Documents to be submitted as an integral part of the offer, and to be evaluated with the tender:

1. Full technical details and catalogues with clearly marked of the type and specifications on of the offered items. 
2. Fully detailed reference list with contact details.

3. Complete Test certificates from a reputable  laboratory to ensure that the required comply with the requirements of this specification.
4. Certificate of calibration according to reputable laboratories shall be attached in addition to that  the thermometer can be user-calibrated.

5. ISO certificate. And any documents that will assist in the evaluation.

· At later stage of evaluation, the most competitive and complied offer shall supply one sample of the above items to be tested for at least four weeks at JEPCO network for approval before awarding the whole quantity.
· Minor deviations of the above specifications might be accepted subject to the approval by the Engineer.   
SCHEDULE (a)
TIME PERIODS FOR PROCUREMENT, DESIGN, MANUFACTURE,
INSPECTION, TESTING, completion AND DELIVERY

Note 1

All time periods are weeks from date of Contract Placement.

Note 2

For purposes of Tender Evaluation,  the time from coomencement date within which the material is required to be delivered to the QAIA shall be 16 weeks.

	Description
	Item1

weeks

	Time within which all arrangement and equipment drawings shall be submitted for approval.
	..........

	Time within which all material shall be available for final inspection and testing in works.
	..........

	Time within which the plant shall be packed and delivered to QAIA.
	..........


schedule (B)

MANUFACTURERS AND PLACES OF MANUFACTURE,
TESTING AND INSPECTION

	Item 

No.
	Description
	Manufacturer
	Place of

Manufacture
	Place of Testing & Inspection

	1.
	Infrared imaging camera
	....................
	....................
	....................


SCHEDULE (C)
 TECHNICAL DATA

· Standard accessories included:

	Item
	Quantity

	
	

	
	

	
	

	
	

	
	

	
	

	
	


· All optional accessories including lenses to be used with the above items shall be quoted separately along with the offer.  
	Item
	Quantity
	Unit Price
	Total price

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	


SCHEDULE (D)
 TECHNICAL DATA

Full Compliance schedule sheet shall be submitted with the offer for each item.

schedule (D1) 
technical schedule

To be completed by the Tenderer
	No
	Description
	Item 1

	1
	Model no.
	
	--------------------------

	2
	Thermal pixels resolution 
	
	--------------------------

	3
	Display Size
	"
	
--------------------------

	4
	Thermal sensitivity 
	mK
	-------------------------

	5
	Temperature Measurement Range
	
	--------------------------

	6
	Accuracy 
	%
	--------------------------

	7
	Frame rate
	Hz
	--------------------------

	8
	Spatial Resolution IFOV
	mrad
	

	9
	a. Standard Field of View
b. Optional Lenses.
	
	--------------------------
--------------------------

	10
	Zoom/ Image Enhancement type
	
	--------------------------

	11
	Focusing system.
	
	--------------------------

	12
	Visible Camera.
	MP
	--------------------------

	13
	Optical Zoom / Telephoto Options.
	
	--------------------------

	14
	Laser pointer 
	Y/N
	--------------------------

	15
	USB interface 
	Y/N
	--------------------------

	16
	a. Wireless connectivity types.

b. GPS.
	Y/N
	--------------------------
--------------------------

	17
	Operating conditions :

a. Humidity.

b. Operating Temperature .
	
	--------------------------
--------------------------

	18
	IP protection without case.
	
	------------------------


schedule (D1) 

technical schedule

To be completed by the Tenderer
	19
	Battery type

Battery capacity

Run time.

Spare battery.
	Y/N
	--------------------------
--------------------------
--------------------------
--------------------------

	20
	Auto power-off feature.

Low battery indication feature.
	Y/N

Y/N
	--------------------------
--------------------------

	21
	Weight
	kg
	--------------------------

	22
	Warranty 
	Years
	--------------------------

	23
	Compliance / Standards
	
	--------------------------

	24
	Training & Support
	
	--------------------------


schedule (E) 

DEVIATIONS FROM SPECIFICATION
schedule (F) 
quantities with prices 

	Description- Item 1
	:
	Infrared imaging camera

	Quantity
	:
	8
	(1)

	Firm unit FOB price 
	:
	

	(2)

	CURRENCY
	:
	

	

	Unit firm freight charges to Amman Airport/Jordan 
	:
	

	(3)

	Unit firm CPT price to Amman Airport/Jordan
	:
	

	(4) = (2) + (3)

	Total firm CPT price to Amman Airport/Jordan
	:
	

	(5) = (1)   (4)


[image: image2.png]



PAGE  
SPECIFICaTION  JEp 24/2026

fp 18-01
         3

